Configuration: Standard lifetime (12.5 ns)

Source: 80 MHz B&H

Signal: SiN emitter (no g2)

Bins: 1024

Sync wire length: X

Result:

Sync wire length: X - 3 feet

Solution: Add ~ 6 feet?

Yes, it worked nearly perfectly.

But it needs to be confirmed that I am referencing off of the correct pulse.

For the CFD, in both cases, -51 mV threshold and -500 mV threshold both work.

Configuration: Standard lifetime (50 ns)

Source: 20 MHz B&H

Signal: SiN emitter (no g2)

Bins: 1024

Sync wire length: X

Result:

The SYNC cable needs to be increased by like 20 feet………. Ruh roh

Added like 12 and it looks great.

Optimal settings

G2 300 ns

Change 6% to 10%

G2 50 ns

Good as is